## explanation
bt | total damping, bo + bc
bo | open-circuit damping | datasheet
bc | electrical damping | calc
Rt | total resistance, Rc + Rload
Rc | feedback coil resistance
Rload | Parallel sum of Rs & Zamp ; = (Rs * Zamp)/(Rs+Zamp) where
Rs | shunt resistor
Zamp | input impedance of preamp
the S & T resistors complicate this - see other spreadsheed calculations
fo | natural freqeuncy (Hz)
m | mass (Kg)
Go | intrinsic sensitivity,or motor constant, sqrt (4*pi*fo*m*Rt*bc)
bc | electrical damping (Go**2)/(4*pi*fo*m*Rt)
poles: two each
p1 | (-2*pi*fo*bt)+j*(2*pi*fo*sqrt(1-bt**2))
p2 | (-2*pi*fo*bt)-j*(2*pi*fo*sqrt(1-bt**2))
zeros: two each
0 | 0
0 | 0
Geff | Effective sensitivity or effective motor constant, Go*(Rload/Rt)

## Qamar 2401a
Seismometer calibration worksheet adapted from Tony Qamar
Basic Geophone calibration procedure, based on Pat McChesney's Notes on Tony Qamar's technique
The goal of this procedure is to calculate the S and T impedance-matching resistors on the McVCO L-pad, to give the desired geophone Motor Constant (G) and damping (B)
1) Choose the desired G. For an L-4, the USGS chooses G=100V/m/sec
2) Choose the desired B. The USGS chooses B=0.8
3) Fill out Table 1 with values from the seismometer
4)
L4c six-pin connector
[9 empty] | pin | color | function
[5 empty] | Table 1 |  |  |  | A | black | coil
[9 empty] | B | red | coil | also case when L4 is on its top or at one position on its side
serial number |  |  |  |  | 2401 |  |  |  | C | green | cal coil
Case to coil resistance | [8 empty] | D | white | cal coil
Seismometer mass(kg) (read from seismometer tag) |  |  |  | M | 0.961 | (mass)
Coil resistance (measure) |  |  |  | rs | 5600 | (coilRes) |  |  | H | shield |  | not connected to case
Measured Damped Period (measure with re open) |  |  |  | Td | 1.1
Choose Ge, the Effective Motor Constant (V/m/sec) |  |  |  | Ge | 100 | (Ge)
Determine Ramp from VCO schematics |  |  |  | Ramp = | 10000 | (Ramp)
Choose desired damping, B |  |  |  | desired B= | 0.8 | (desiredB)
[5 empty] | e=A1/A2 | del=ln(e) | B2=del2/(del2+4n2pi2) | Tn=Td*sqrt(1-B02) | gamma=(rs+re)*(B-B0) |  |  | G=sqrt(4*pi*mass*gamma/Tn)
|  | volts | volts
| re | A1 | A2 | n | e | del | B (B0 if re is open) | Natural Free Period, Tn (NatlFreePer) | gamma |  |  | G
| 1000000000 | 2.9 | 1.16 | 0.5 | 2.5 | 0.91629 | 0.28157 | 1.0555 | 0 |  |  | 0
| 100000 | 2.6 | 0.83 | 0.5 | 3.1325 | 1.1418 | 0.34344 |  | 6532.8797 |  |  | 273.3261
| 50000 | 2.15 | 0.51 | 0.5 | 4.2157 | 1.4388 | 0.4185 |  | 7612.8626 |  |  | 295.0549
| 33000 | 1.94 | 0.38 | 0.5 | 5.1053 | 1.6303 | 0.46282 |  | 6996.0982 |  |  | 282.8504
| 25000 | 0.692 | 0.108 | 0.5 | 6.4074 | 1.8575 | 0.51124 |  | 7027.9062 |  |  | 283.4926
[10 empty] | 7042.4367 | (GammaAvg) |  | 283.681 | (G)
calculate S & T resistors for desired damping and motor constant
|  |  |  | re=(gammaavg/(B-B0)) - rs |  | re= | 7984.2863 | (Re)
|  |  |  | rab=(Ge/G)*(re+rs) |  | rab= | 4788.5783 | (Rab)
|  |  |  | S=1/((1/rab)-(1/ramp)) |  | S= | 9188.6216 |  | 7500
|  |  |  | T=re-rab |  | T= | 3195.708 |  | 2200

## L4 puti yyyymmdd
L-4 Seismometer calibration worksheet adapted from Tony Qamar
Use this sheet as a permanent record of seismometer calibration by changing the name to the serial number and the date of the calibration.
Refer to the Procedure, or the document 'seismometer calibration for CVGHM.docx' for instructions
italicised names refer to excel cell names used in calculations on this worksheet
Table 1
serial number |  |  |  |  | test
Case to coil resistance
Seismometer mass (kg) (read from seismometer tag) ( ~1.0) |  |  |  | M |  | (mass)
measure Coil resistance(ohms) ( Measure!) may be 510 |  |  |  | rs |  | (coilRes)
Measured Damped Period (sec) (measure with re open) (~1.0) |  |  |  | Td |  | (t_d)
Choose Ge, the Effective Motor Constant (V/m/sec) |  |  |  | Ge | 100 | (Ge)
Determine Ramp from VCO schematics |  |  |  | Ramp = | 10000 | (Ramp)
Choose desired damping, B |  |  |  | desired B= | 0.8 | (desiredB)
|  |  |  | B12=del2/(del2+pi2) | Tn=Td*sqrt(1-B02) | gamma=(rs+re)*(B-B0) |  | G=sqrt(4*pi*mass*gamma/Tn)
| ohms | volts | volts
| re | A1 | A2 | B1 (B0 if re is open) | Natural Free Period, Tn (NatlFreePer) | gamma |  | G
| 1000000000 |  |  | #DIV/0! | #DIV/0!
| 100000 |  |  | #DIV/0! |  | #DIV/0! |  | #DIV/0!
| 50000 |  |  | #DIV/0! |  | #DIV/0! |  | #DIV/0!
| 33000 |  |  | #DIV/0! |  | #DIV/0! |  | #DIV/0!
| 25000 |  |  | #DIV/0! |  | #DIV/0! |  | #DIV/0!
[6 empty] | #DIV/0! | (GammaAvg) | #DIV/0! | (G)
calculate S & T resistors for desired damping and motor constant
|  |  |  | re=(gammaavg/(B-B0)) - rs |  | re= | #DIV/0! | (Re)
|  |  |  | rab=(Ge/G)*(re+rs) |  | rab= | #DIV/0! | (Rab)
VDAP values
|  |  |  | S=1/((1/rab)-(1/ramp)) |  | S= | #DIV/0! |  | 7500
|  |  |  | T=re-rab |  | T= | #DIV/0! |  | 2200
cross-check on Ge and B for S & T values for this seismometer
| Serial Number | Rs | Rt (=T1+T2) | Ge v/m/sec | B0 | B1 | B
| test | 8330 | 3348 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!

## seisan inputs, PS-2, PSN
values to calculate digital vs PS-2 comparison | Digital | PS-2
Choose output format: | 1 | 1 | 1 | seisan FAP
type of sensor | 2 | 2 | 2 | seismometer
seismometer natural period | 1 | 1 | 1 | for L-4 1-hz seismometers
|  |  | 0.5 | for L-22 2-hz seismometers (3-comps)
seismometer damping ratio | 0.8 | 0.8 | 0.8
sensor loaded generator constant (V/M/S or V/G) | 100 | 100 | 100 V/M/S | for L-4 1-hz seismometers
|  |  | 25 V/M/S | for L-22 2-hz seismometers (3-comps)
|  |  | 10 V/M/S | for BPPTK 10(?)-hz seismometer
recording media gain (count/volt or m/v) |  |  | 819.6 | 12-bit AD conversion: 0.00122v/ct , or 819.6 ct/V
| 6553.6 |  | 6553.6 | psn 16-bit over +/-5V range
|  |  | 3276.8 | Guralp 16-8R over +/-10V range
PS-2 pen sensitivity mm/mV (switch 1,2,3,4)
|  |  |  | one | two | three | four
PS-2 pen sensitivity |  | 0.0078 |  | 0.25 | 0.0156 | 0.0078 | 0.0039
amplifier gain (dB) |  |  | 55 | for standard VDAP installation with 60dB McVCO gain |  |  |  | L4
Digital system: McVCO gain -5dB to account for McVCO/Mc8 conversion mis-match in USGS standard | 61 | 67 | 61 | for standard VDAP installation with 66dB McVCO gain
PS-2: McVCO gain +1.1dB for McVCO/PS-2 conversion mismatch |  |  | 67 | for standard VDAP installation with 72dB McVCO gain |  |  |  | L22
Mc8 output is +/-2.5V / +/-125hz
McVCO output is 115hz / 4.05V
USGS McVco out = 28.4 Hz/V
USGS discrim out=0.02V/Hz
Kinemetrics discrim out =? 0.04 V/Hz |  |  | 1.1 | 20log1.136 | McVCO, Kinemetrics
PS-2 Atelem = McVCO* Disc= 1.136 |  |  | -4.9 | 20log 0.568 | McVCO, Mc8
Digital Atelem=vco * discrim =0.568
|  | 31 |  | PS-2 gain switch (0, 12, 18, 24, 30, 36, 42, 48, 54, 60, 66, 72)
number of filters | 2 | 3 | 2 | 2, for digital system; 3, for McVCO/PS-2
frequency and poles for each filter, poles negative for high pass | 30 4 | 30 4 | 30 4 | McVCO low-pass filter: 4 poles at 30Hz
| 20 4 |  | 20 4 | J120D discriminator low-pass filter: 4 poles at 20Hz
|  | 12.5 2 |  | PS-2 low-pass switch 12dB/octave
|  | 0.1, -5 | 0.1, -5 | PS-2 high pass, 12bB/octave, on Amp board, f=DC, 0.1, 5
type of response |  |  | 1 | displacement
file name for files with poles and zeros, return for no file |  |  | <return>
file name for tabulated values, return for no file |  |  | <return>
give file name for measured values, return for none |  |  | <return>
PROGRAM NOW GENERATES ON-SCREEN GRAPHS
SEISAN rsponse file (y/n=default) |  |  | y
enter station code |  |  |  | enter station name
enter component (4 chars0 e.g. SH Z) |  |  | SH Z | if short-period vertical
enter date as yyyymmddhhmmss at least up to the day (e.g. 19880123)
Latitude, return for none |  |  | <return>
Longitude, return for none
elevation in meters, return for none
enter comments, one line |  |  | VDAP installation yadda yadda

## SP damping
T(f) =((f^2)/((((f0^2-f^2)^2)+4*(damping^2)*(f0^2*f^2))^0.5)*motorConst
|  |  | G | Seismometer motor constant
|  |  | Rs | parallel (shunt) resistance | [13 empty] | T(f) =((f^2)/(f0^2-f^2+2*damping*f0*f)
|  |  | Rt | series (t) resistance | [5 empty] | B=B0+B1
|  |  | Rseis | seismometer resistance | [5 empty] | B(0) | mechanical damping of seismometer
|  |  | Ramp | VCO input inpedance | [5 empty] | B(1) | electromagnetic damping | [5 empty] | poles
[17 empty] | pole 1 |  | pole 2
| tabla de calculos de damping, dado ciertos resistencias y valores estandardes de otros parametros | [15 empty] | (-2*pi*fo*b)+j*(2*pi*fo*sqrt(1-b**2)) |  | (-2*pi*fo*b)-j*(2*pi*fo*sqrt(1-b**2))
sensor | G (v/cm/sec) | f0 (Hz) | Rs | Rt (=T1+T2) | Rseis | Ramp | true Rs: sum of Rs & Ramp | Geff v/cm/sec | B(0) | m kg | B(1) | B | Bsercel | T(f) (sec-1) | T(f) (sec-1) | T(w) (sec^-1_ | (-2*pi*fo*b) | +j*(2*pi*fo*sqrt(1-b**2) | (-2*pi*fo*b) | -j*(2*pi*fo*sqrt(1-b**2)
Trillium | 7.5 | 0.0083 | 2000 | 0 | 100000 | 10000 | 1666.6667 | 0.12295 | 0.28 | 1 | 53.048 | 53.328 | 1.362 | 9.2065 | -107.2058 | -8.7194 | -2.781 | #NUM! | -2.781 | #NUM!
L4 | 2.8 | 1 | 8890 | 0 | 5500 | 100000000 | 8889.2097 | 1.7298 | 0.28 | 1 | 0.43359 | 0.71359 | 0.70045 | 121.2003 | 121.2003 | 0.70068 | -4.4835 | 4.4016 | -4.4835 | -4.4016 | USGS standard Rs & Rt values
| 2.84 | 1 | 9189 | 3196 | 5600 | 10000 | 4788.681 | 1.0011 | 0.28 | 0.961 | 0.49166 | 0.77166 | 0.87295 |  |  |  | -4.8484 | 3.9962 | -4.8484 | -3.9962 | from Qamar 2401a calculations
| 2.8 | 1 | 6200 | 2300 | 5500 | 100000 | 5838.0414 | 1.1986 | 0.28 | 1 | 0.45747 | 0.73747 | 0.8136 |  |  |  | -4.6336 | 4.2434 | -4.6336 | -4.2434 | Rs & Rt values provided to Carlos by persons unknown
| 2.8 | 1 | 13000 | 0 | 5500 | 100000 | 11504.4248 | 1.8944 | 0.28 | 1 | 0.36691 | 0.64691 | 0.63579 |  |  |  | -4.0645 | 4.7912 | -4.0645 | -4.7912 | sercel-kinemetrics values for 0.6 damping (underdamped)
| 2.8 | 1 | 7500 | 2200 | 5500 | 100000 | 6976.7442 | 1.331 | 0.28 | 1 | 0.4251 | 0.7051 | 0.7649 |  |  |  | -4.4301 | 4.4553 | -4.4301 | -4.4553 | USGS Rs, Rt values, Rsamp = 100,000
| 2.8 | 1 | 1000000000 | 0 | 5500 | 100000 | 99990.001 | 2.654 | 0.28 | 1 | 0.059144 | 0.33914 | 0.33735 |  |  |  | -2.1308 | 5.9106 | -2.1308 | -5.9106 | damping of Eentec with no resistors added; seriously underdamped
| 2.8 | 1 | 6134 | 0 | 5500 | 100000 | 5779.4863 | 1.4347 | 0.28 | 1 | 0.55313 | 0.83313 | 0.81637 |  |  |  | -5.2346 | 3.475 | -5.2346 | -3.475 | sercel/kinemetrics values for 0.8 damping (slightly overdamped)
| 2.8 | 1 | 20000 | 0 | 5500 | 10000 | 6666.6667 | 1.5342 | 0.28 | 1 | 0.5128 | 0.7928 | 0.77726 |  |  |  | -4.9812 | 3.8294 | -4.9812 | -3.8294
| 2.8 | 1 | 6200 | 5000 | 5500 | 100000 | 5838.0414 | 1.0005 | 0.28 | 1 | 0.38187 | 0.66187 | 0.8136 |  |  |  | -4.1585 | 4.7098 | -4.1585 | -4.7098 | USGS values if Ramp left at 100,000
| 2.8 | 1 | 6200 | 4900 | 5500 | 100000 | 5838.0414 | 1.0067 | 0.28 | 1 | 0.38422 | 0.66422 | 0.8136 |  |  |  | -4.1733 | 4.6967 | -4.1733 | -4.6967 | USGS values if Ramp left at 100,000
[17 empty] | 0 | 0 | 0 | 0
L22D | 1.0961 | 2 | 7875 | 9869 | 5470 | 10000 | 4405.5944 | 0.24457 | 0.46 | 0.0728 | 0.33257 | 0.79257 | 0.80076 |  |  |  | -9.9594 | 7.6626 | -9.9594 | -7.6626 | L22d. 5470 R(S) using Pats Rs & Rt values Sept 2008
| 1.0961 | 2 | 12100 | 8040 | 5470 | 10000 | 5475.1131 | 0.3161 | 0.46 | 0.0728 | 0.34587 | 0.80587 | 0.81439 |  |  |  | -10.1266 | 7.4402 | -10.1266 | -7.4402 | L22d. 5470 R(S) using standard VDAP Rs & Rt values Sept 2008
| 0.33468 | 2 | 12100 | 8040 | 510 | 10000 | 5475.1131 | 0.13065 | 0.46 | 0.0728 | 0.043652 | 0.50365 | 0.50473 |  |  |  | -6.3289 | 10.8559 | -6.3289 | -10.8559 | L22d. 510ohm using standard VDAP Rs & Rt values
| 1.0961 | 2 | 1000000 | 7000 | 5470 | 10000 | 9900.9901 | 0.4851 | 0.46 | 0.0728 | 0.29352 | 0.75352 | 0.76075 |  |  |  | -9.4688 | 8.2612 | -9.4688 | -8.2612
| 0.33599 | 2 | 10000000 | 45000 | 514 | 10000 | 9990.01 | 0.060474 | 0.72 | 0.0728 | 0.011117 | 0.73112 | 0.73139 |  |  |  | -9.1872 | 8.5732 | -9.1872 | -8.5732
L22E | 0.41759 | 2 | 67000 | 8000 | 519 | 2000000000 | 66997.7556 | 0.37048 | 0.72 | 0.0728 | 0.012621 | 0.73262 | 0.72845 |  |  |  | -9.2061 | 8.5529 | -9.2061 | -8.5529 | L22E at newberry, McRick instrumentation amp
L22E | 0.41759 | 2 | 100000000 | 0 | 519 | 10000 | 9999.0001 | 0.39698 | 0.72 | 0.0728 | 0.090615 | 0.81062 | 0.78069 |  |  |  | -10.1862 | 7.3584 | -10.1862 | -7.3584 | L22E at newberry, McRick instrumentation amp
| 0.41759 | 2 | 12100 | 8040 | 519 | 10000 | 5475.1131 | 0.16291 | 0.72 | 0.0728 | 0.067912 | 0.78791 | 0.76549 |  |  |  | -9.9009 | 7.738 | -9.9009 | -7.738 | L22e. 510ohm using standard VDAP Rs & Rt values
873-vert | 0.41436 | 2 | 67000 | 8000 | 511 | 2000000000 | 66997.7556 | 0.36765 | 0.72 | 0.0728 | 0.012428 | 0.73243 | 0.72832 |  |  |  | -9.2037 | 8.5555 | -9.2037 | -8.5555 | L22E at newberry, McRick instrumentation amp
873-trans | 0.41678 | 2 | 67000 | 8000 | 517 | 2000000000 | 66997.7556 | 0.36977 | 0.72 | 0.0728 | 0.012573 | 0.73257 | 0.72842 |  |  |  | -9.2055 | 8.5535 | -9.2055 | -8.5535 | L22E at newberry, McRick instrumentation amp
874-long | 0.41638 | 2 | 67000 | 8000 | 516 | 2000000000 | 66997.7556 | 0.36942 | 0.72 | 0.0728 | 0.012548 | 0.73255 | 0.7284 |  |  |  | -9.2052 | 8.5539 | -9.2052 | -8.5539 | L22E at newberry, McRick instrumentation amp
874-vert | 0.41517 | 2 | 67000 | 8000 | 513 | 2000000000 | 66997.7556 | 0.36836 | 0.72 | 0.0728 | 0.012476 | 0.73248 | 0.72836 |  |  |  | -9.2043 | 8.5548 | -9.2043 | -8.5548 | L22E at newberry, McRick instrumentation amp
874-trans | 0.41476 | 2 | 67000 | 8000 | 512 | 2000000000 | 66997.7556 | 0.36801 | 0.72 | 0.0728 | 0.012452 | 0.73245 | 0.72834 |  |  |  | -9.204 | 8.5552 | -9.204 | -8.5552 | L22E at newberry, McRick instrumentation amp
\
GS-20 | 0.28 | 10 | 7500 | 2200 | 400 | 10000 | 4285.7143 | 0.17427 | 0.3 | 0.011 | 0.082372 | 0.38237 | 0.3939 |  |  |  | -24.0244 | 58.0555 | -24.0244 | -58.0555 | BPPTK GS-20DX, 10Hz, from datasheet, L4 resistors
| 0.28 | 10 | 10000000 | 1000 | 400 | 10000 | 9990.01 | 0.24558 | 0.3 | 0.011 | 0.049797 | 0.3498 | 0.34235 |  |  |  | -21.9777 | 58.8608 | -21.9777 | -58.8608 | BPPTK GS-20DX, 10Hz, from datasheet
| 0.28 | 10 | 10000000 | 0 | 400 | 10000 | 9990.01 | 0.26922 | 0.3 | 0.011 | 0.05459 | 0.35459 | 0.34235 |  |  |  | -22.2789 | 58.7474 | -22.2789 | -58.7474 | BPPTK GS-20DX, 10Hz, into 10,000 ohm PSN board
| 0.28 | 10 | 500 | 0 | 400 | 10000 | 476.1905 | 0.15217 | 0.3 | 0.011 | 0.64733 | 0.94733 | 0.80217 |  |  |  | -59.5209 | 20.1213 | -59.5209 | -20.1213 | BPPTK GS-20DX, 10Hz, from datasheet
| 0.28 | 10 | 1000 | 0 | 400 | 10000 | 909.0909 | 0.19444 | 0.3 | 0.011 | 0.43327 | 0.73327 | 0.63611 |  |  |  | -46.0712 | 42.7206 | -46.0712 | -42.7206 | BPPTK GS-20DX, 10Hz, 1k damping resistor placed across amplifier inputs
L10AR | 0.31 | 10 | 1000000000 | 0 | 374 | 1000000000 | 500000000 | 0.31 | 0.3 | 0.0122 | 1.2537e-06 | 0.3 | 0.3
[7 empty] | Rc | sqrt (Rc) | v/inch/sec | v/m/sec | v/cm/sec
| L4 transduction =7.0v/inch/sec | [7 empty] | 7 | 273 | 2.8
| L22D transduction 510 ohm coil = (.038)*(sqrt Rc) |  |  |  |  | 0.038 | 510 | 22.5832 | 0.85816 | 33.4683 | 0.33468
| L22E transduction 510 ohm coil |  |  |  |  | 0.047 | 510 | 22.5832 | 1.0614 | 41.395 | 0.41395
| GS-20DX | [7 empty] | 0.7 | 27.3 | 0.273
| L10AR transduction = 0.041*(sqrt374) |  |  |  |  | 0.041 | 374 | 19.3391 | 0.7929 | 30.9232 | 0.30923
| parallel resistance |  |  | R=1/((1/R1)+(1/R2)) |  | r1 | r2 | R
[6 empty] | 1 | 2 | 0.66667
[6 empty] | 1000 | 400 | 285.7143
[6 empty] | 280 | 10000 | 272.3735
digitizer sensitivity
|  | device | range V | bits | counts | ct/v
|  | PSN ADC | 10 | 16 | 65536 | 6553.6
|  | Guralp 16R | 20 | 16 | 65536 | 3276.8
note 1 | Damping B=B0 +B1; mechanical (B0) + electromagnetic (B1). Mechanical damping es fijo por el deseno en 0.28. electromagnetico se puede ajustar por anadir resistencias
Damping varia desde 0 hasta 1. 'Critical damping' =0.707 (1/rais de 2) 0 3dB por abajo del maximo.
Estandar USGS =0.8 damping.
El Eentec, sin resistentes adicionales, B= solo mechanical damping, 0.3. entonces muy baja damping
nota 2 | G= transductance=motor constant=generator constant en unidades de v/cm/seg
Geff =G effectivo, dado reducion en el por factores de los resistencias
G del L-4= 2.8v/cm/seg
el Estandar USGS es G=1v/cm/seg.
nota 3 | Ramp = impedencia de la entrada al amplificador. Todos los USGS VCO, McVCO Ramp=10,000.
Eentec Ramp=100,000
el valor de Ramp afecta B y Geff.
Sugiero uno de dos modificaciones.
Mod 1: solder una resistancia de 6134 cruzando el resistencia R301, el 100,000 ohmio Ramp.
Esta da un B=0.8 y Geff=1.4v/cm/seg
Mod 2. hacer Mod 1, y anadir una resistencia de 5000 en serie con el inductor L301
Esta da un B=0.8 y Geff=1.0v/cm/seg
Note on K9: Andy Lockhart: Sercel/kinemetrics: damping is b0 + b1 where b1=1.1(rseis)/(Rseis+Rs)
Note on H14: Andy Lockhart: the true Rs, when Ramp is taken into account. This is the value to compare to the sercel-kinemetrics Rs values.
Note on N14: Andy Lockhart: Sercel/kinemetrics: damping is b0 + b1 where b1=1.1(rseis)/(Rseis+Rs) This assumes Rt=0 so it gives incorrect results when Rt not 0.
Note on O14: Lockhart, Andrew B.: my formulation. Not used in 2021 document Just a chek on the 2021 formula from Jens
Note on D23: Andy Lockhart: valor muy alta para eliminar Rs del calculo
Note on A29: Max: From Sercel datasheet, G=0.038(sqrt(Rc))v/inch/sec. for a L22D
Note on N29: Andy Lockhart: this includes term for RT since the RS term in sercel includes all the exernal resistance. It also uses the 1.23 constant for the L22D.
Note on E30: Andy Lockhart: Rt=2xRtresistor values; one Rt on each leg doubles effective Rt.
Note on N30: Andy Lockhart: this includes term for RT since the RS term in sercel includes all the exernal resistance. It also uses the 1.23 constant for the L22D.
Note on E31: Andy Lockhart: Rt=2xRtresistor values; one Rt on each leg doubles effective Rt.
Note on N31: Andy Lockhart: this includes term for RT since the RS term in sercel includes all the exernal resistance. It also uses the 1.23 constant for the L22D.
Note on E32: Andy Lockhart: Rt=2xRtresistor values; one Rt on each leg doubles effective Rt.
Note on N32: Andy Lockhart: this includes term for RT since the RS term in sercel includes all the exernal resistance. It also uses the 1.23 constant for the L22D.
Note on E33: Andy Lockhart: Rt=2xRtresistor values; one Rt on each leg doubles effective Rt.
Note on N33: Andy Lockhart: this includes term for RT since the RS term in sercel includes all the exernal resistance. It also uses the 1.23 constant for the L22D.
Note on A34: Max: From Sercel datasheet, G=0.047(sqrt(Rc))v/inch/sec. for a L22E
Note on E34: Andy Lockhart: Rt=2xRtresistor values; one Rt on each leg doubles effective Rt.
Note on N34: Andy Lockhart: this includes term for RT since the RS term in sercel includes all the exernal resistance. It also uses the 1.23 constant for the L22D.
Note on A35: Max: From Sercel datasheet, G=0.047(sqrt(Rc))v/inch/sec. for a L22E
Note on E35: Andy Lockhart: Rt=2xRtresistor values; one Rt on each leg doubles effective Rt.
Note on N35: Andy Lockhart: this includes term for RT since the RS term in sercel includes all the exernal resistance. It also uses the 1.23 constant for the L22D.
Note on E36: Andy Lockhart: Rt=2xRtresistor values; one Rt on each leg doubles effective Rt.
Note on N36: Andy Lockhart: this includes term for RT since the RS term in sercel includes all the exernal resistance. It also uses the 1.23 constant for the L22D.
Note on E37: Andy Lockhart: Rt=2xRtresistor values; one Rt on each leg doubles effective Rt.
Note on N37: Andy Lockhart: this includes term for RT since the RS term in sercel includes all the exernal resistance. It also uses the 1.23 constant for the L22D.
Note on E38: Andy Lockhart: Rt=2xRtresistor values; one Rt on each leg doubles effective Rt.
Note on N38: Andy Lockhart: this includes term for RT since the RS term in sercel includes all the exernal resistance. It also uses the 1.23 constant for the L22D.
Note on E39: Andy Lockhart: Rt=2xRtresistor values; one Rt on each leg doubles effective Rt.
Note on N39: Andy Lockhart: this includes term for RT since the RS term in sercel includes all the exernal resistance. It also uses the 1.23 constant for the L22D.
Note on E40: Andy Lockhart: Rt=2xRtresistor values; one Rt on each leg doubles effective Rt.
Note on N40: Andy Lockhart: this includes term for RT since the RS term in sercel includes all the exernal resistance. It also uses the 1.23 constant for the L22D.
Note on E41: Andy Lockhart: Rt=2xRtresistor values; one Rt on each leg doubles effective Rt.
Note on N41: Andy Lockhart: this includes term for RT since the RS term in sercel includes all the exernal resistance. It also uses the 1.23 constant for the L22D.

## full eq
This is a full description of the analog system response
|  | Seismometer |  |  |  | McVCO |  |  |  |  | Discriminator |  | ADC
diagram |  | m/sec - V conversion |  | L-pad |  | V-Hz conversion |  | Gain | 30-hz lo pass filter |  | Hz-V conversion | 20Hz lo pass filter | V-ct conversion
Transfer function T(f) |  |  | T(f) =((f^2)/((((f0-f^2)^2)+4*(damping^2)*(f0^2)*(f^2))^0.5)*MotorConst)*(10^(GaindB/20))*(1/(1+(f/30)^8)^0.5)*(1/(1+(f/20)^8)^0.5)*ADC
where | f | 1 | hz | frequency (hz)
| f0 | 1 | hz | seismometer center frequency, Hz, (1Hz for an L4)
| Damping, B | 0.8 | v/v | damping, set by seismometer and L-pad, nominally 0.8
| MotorConst | 100 | v/m/sec | motor constant, set by seismometer and L-pad, nominally 100V/m/sec
| McVCOgain_dB | 72 | dB | Gain of McVCO, dB, =10^(gain/20) 60dB=1000, 66dB=1995, 72dB= 3981
| McVCOloPass | 30 | hz | McVCO lo-pass filter corner freq
| McVCO_V_Hz | 28.3951 | Hz | McVCO output is 115hz / 4.05V
| Disc_Hz_V | 0.02 | V | Discrim output is +/-2.5V / +/-125hz
| DiscLoPass | 20 | Hz | Disc lo-pass corner freq
| Mc8HiPass | 0.1 | Hz | Mc8 hi-pass filter
| ADC | 3276.8 | ct/V | analog-digital converter sensitivity; 16-bit Guralp ADC, 20V range
seismometer transfer function, Poles & zeros, V/m/sec
62.5 | T(f) seismometer&L-pad, V/m/sec |  |  |  |  | T(f) =((f^2)/((((f0^2-f^2)^2)+4*(damping^2)*(f0^2*f^2))^0.5)*motorConst
0.625 | [5 empty] | T(f) =((f^2)/((((f0^2-f^2)^2)+4*(damping^2)*(f0^2*f^2))^0.5)
0.625 | [5 empty] | T(f) =((f^2)/(f0^2-f^2+2*damping*f0*f)
| (-2*pi*fo*b)+j*(2*pi*fo*sqrt(1-b**2)) | [5 empty] | (-2*pi*fo*b)-j*(2*pi*fo*sqrt(1-b**2))
2 poles | (-2*pi*fo*b) |  | +j*(2*pi*fo*sqrt(1-b**2) |  |  |  | (-2*pi*fo*b) |  | -j*(2*pi*fo*sqrt(1-b**2)
| -5.0265 |  | 3.7699 |  |  |  | -5.0265 |  | -3.7699
2 zeros | 0 |  | 0 |  |  |  | 0 |  | 0
Note: to convert the velocity transfer function to displacment, multiply velocity by (2*3.1415*f)
McVCO transfer function, Gain, filter, Hz/V conversion
3981.0717 | adjustable McVCO gain, 10^(McVCOgain_dB/20) |  |  |  |  | 3981 | 72dB assumption)
1 | McVCOloPass 30Hz 4-pole filter |  |  |  |  | McVCOlopass=(1/((1+(B34/D34)^8)^0.5))
1262272898.2161 | PAZ constant, w^4 | [5 empty] | -b+/- ((b^2-4*a*c)^.5)/2*a
[5 empty] | f0(hz) | f0(w) | a | b | c | p1 | p2
|  | (s^2 + 0.765367*s*wc +wc^2) |  |  | 30 | 188.4954 | 1 | 144.2682 | 35530.5158 | -72.1341 | 174.147
[10 empty] | -72.1341 | -174.147
|  | (s^2 + 1.8477759*s*wc +wc*2) |  |  | 30 | 188.4954 | 1 | 348.2973 | 35530.5158 | -174.1486 | 72.1302
[10 empty] | -174.1486 | -72.1302
28.3951 | McVCO output is 115hz / 4.05V
Discriminator transfer function: filter, hz to V conversion
0.02 | Discrim output is +/-2.5V / +/-125hz
Note mismatch of McVCO Hz to V and Discrim V to Hz conversions. Ideally this should equal 1 but the VDAP McVCO and the discriminators have different conversion factors due to historical reasons. VDAP McVCO uses the USGS standard of115Hz/4.05V, dating from when USGS McVCOs were powered by 4.05V mercury cells whose output voltage is stable enough for the VCO design to use it as a voltage standard.
1 | Disc 20Hz lo-pass 4th order filter
-b+/- ((b^2-4*a*c)^.5)/2*a
249337856.4377 | PAZ constant, w^4 |  |  |  | f(hz) | f(w) | a | b | c | p1 | p2
|  | (s^2 + 0.765367*s*wc +wc^2) |  |  | 20 | 125.6636 | 1 | 96.1788 | 15791.3404 | -48.0894 | 116.098
[10 empty] | -48.0894 | -116.098
|  | (s^2 + 1.8477759*s*wc +wc*2) |  |  | 20 | 125.6636 | 1 | 232.1982 | 15791.3404 | -116.0991 | 48.0868
[10 empty] | -116.0991 | -48.0868
ADC
3276.8
| total gain outboard of seismometer at 1 hz |  |  |  |  | normalization factor
3981.0717 | McVCOgain_dB |  |  |  |  | 1 | v/m/sec seismometer normalization factor
1 | McVCO lo pass ^4 (approx)
28.3951 | McVCO_V_Hz | -0.5679 |  |  |  | 1262000000 | McVCO lo-pss 30z
0.02 | Disc_Hz_V |  |  |  |  | 249400000 | disc lo-pass 20hz
1 | Disc lo pass ^4 (approx) |  |  |  |  | 3.1474e+17 | total normalization factor after seismometer
67.0855 | total gain dB (McVCO gain * telem attenuation) |  |  |  |  | was 67dB
3276.8 | ADC
7408371.4218 | total gain of 'datalogger' in stationXML
62.5 | T(1hz) | gain of 'sensor' in stationXML
463023213.861 | ct/m/sec | total system gain in stationXML. This checks with the total gain at 1Hz calculated in GSE2 response file
PAZ for the whole system, in one line, for GEMPA input
-5.0265484 +3.7699113i , -5.0265484 -3.7699113i , -72.134140 +174.14719i , -72.134140 -174.14719i , -174.14720 +72.134102i , -174.14720 +72.134102i , -48.089432 +116.09813i , -48.089432 -116.09813i, -116.09813 +48.08942i , -116.09813 -48.08942i
| 188.49 | w(30Hz)
| 125.66 | w(20Hz)

## scratch sheet
review of Sulis' 8/12 'update on stationXML.doc'
| nm/sec | m/sec
seiscomp max amp | 8284 | 8.284e-06
seisan max amp | 6279 | 6.279e-06
2005
0.24203
cal pulse | 100 | 101011101001 | 1011 | '0000' | '0110'
gain | 66

## ct_v
Bit | counts | bit % | power dB | voltage dB
0 | 1 | 100 | 0 | 0
1 | 2 | 50 | 3.0103 | 6.0206
2 | 4 | 25 | 6.0206 | 12.0412
3 | 8 | 12.5 | 9.0309 | 18.0618
4 | 16 | 6.25 | 12.0412 | 24.0824
5 | 32 | 3.125 | 15.0515 | 30.103
6 | 64 | 1.5625 | 18.0618 | 36.1236
7 | 128 | 0.78125 | 21.0721 | 42.1442
8 | 256 | 0.39062 | 24.0824 | 48.1648
9 | 512 | 0.19531 | 27.0927 | 54.1854
10 | 1024 | 0.097656 | 30.103 | 60.206
11 | 2048 | 0.048828 | 33.1133 | 66.2266
12 | 4096 | 0.024414 | 36.1236 | 72.2472 | PCI_MIO_16E_4 | .00122V/BIT | +/-5v range
13 | 8192 | 0.012207 | 39.1339 | 78.2678
14 | 16384 | 0.0061035 | 42.1442 | 84.2884
15 | 32768 | 0.0030518 | 45.1545 | 90.309
16 | 65536 | 0.0015259 | 48.1648 | 96.3296 | PSN
17 | 131072 | 0.00076294 | 51.1751 | 102.3502
18 | 262144 | 0.00038147 | 54.1854 | 108.3708
19 | 524288 | 0.00019073 | 57.1957 | 114.3914
20 | 1048576 | 9.5367e-05 | 60.206 | 120.412
21 | 2097152 | 4.7684e-05 | 63.2163 | 126.4326
22 | 4194304 | 2.3842e-05 | 66.2266 | 132.4532
23 | 8388608 | 1.1921e-05 | 69.2369 | 138.4738
24 | 16777216 | 5.9605e-06 | 72.2472 | 144.4944
|  |  |  | gain | dB
note: log10=1 |  |  |  | 5500 | 74.8073
|  |  |  | 1000 | 60
Shortcuts for dB: Voltage - |  |  |  | 1700 | 64.609
| double voltage (x2), add 6dB. |  |  | 6000 | 75.563
| X10 voltage, add 20dB |  | 60dB
for NI ADC in EW system: | 12-bit counts | volts | volts / count | gain = 2 |  |  | worm system - disc out is +/-2.5vpp, gain =4 on NI software for +/-10vpp or 12dB
| 4096 | 10 | 0.0024414 | 0.0012207 | =1.22mV/bit |  | signal saturates at +/-2048 for +/-2.5V
| 4096 | 10 | 0.0024414 | 0.0012207
| counts | volts | counts/volt | volts/count |  |  | worm system - disc out is +/-2.5Vpp, PSN range is 10V,
psn 16-bit board | 65536 | 10 | 6553.6 | 0.00015259 |  |  | signal saturates at 16k
guralp 16-bit DM16-R8 | counts | volts | counts/volt | volts/count |  | saturates at 5V
| 65536 | 20 | 3276.8 | 0.00030518 |  | 16384
| 65536 | 10 | 6553.6